# Annual Inspection Checklist for a PC3 Facility

# Checklist for annual inspection against the usual Conditions of Certification as detailed in the Gene Technology Regulator’s Guidelines for Certification of a Physical Containment Level 3 Facility Version 1.0 - 01 December 2022

| # Name(s) and signature(s) of person(s) inspecting the facility (please print name clearly) |  |
| --- | --- |

| # Date of Inspection |  |
| --- | --- |

Please Note

- The use of this checklist proforma is not mandatory in order to satisfy the annual inspection reporting component of Condition 15. Rather, it is provided to assist those who find it convenient to use in the annual inspection of certified facilities for compliance with the conditions of certification under Condition 15.

- A completed copy of this proforma will be accepted by the OGTR as the annual inspection report for a certified facility under Condition 15, but the proforma is not intended to be the only acceptable format for the report.

- Please use the ‘Application Checklist’ against the requirements for certification (as opposed to this ‘Annual Inspection Checklist’) when applying for a new certification, or when seeking a variation to the requirements for certification of a facility (e.g. lifting the suspension of a certification after modifications to the facility.)

- Please do not send this report to the OGTR unless specifically requested.

# Facility details

| Organisation |  |
| --- | --- |
| Facility details | Cert-xxx PC3 Facility ☐ Animal room(s) ☐ Invertebrates room(s) |
| GMOs in the facility |  |
| Relevant non-GMOs present |  |
| Licences associated with facility |  |
| NLRD(s) associated with facility |  |

# Facility management

## Facility manual

Person who is responsible for training of staff: __________________________________________________

Date of last review: ________________________________________________________________________

Information required in the facility manual:

| Required information | Addressed |
| --- | --- |
| Facility manager’s contact details | ☐ |
| Emergency contacts details | ☐ |
| Copy of certification instruments and conditions of certification | ☐ |
| Copy of licences/NLRDs | ☐ |
| Copy of Guidelines for the Transport, Storage and Disposal of GMOs & Certification of a PC3 facility | ☐ |
| Structure & operation of facility (including design limits) | ☐ |
| Details of all GM and non-GM organisms used in the facility, associated risks, and risk management strategies | ☐ |

Procedures to be included in the facility manual:

| Procedures | Addressed |
| --- | --- |
| Procedures for handling animals and invertebrates held within the facility, if applicable, including the procedure to be implemented in the event an animal/invertebrate escape within or outside the facility and to account of all animal/invertebrate held within the facility | ☐ |
| Entry & exit procedures, PPE requirements and protocols for use / removal | ☐ |
| Procedures for use and operation of BSC / other aerosol containment equipment, including the removal of gloves and over-sleeves on completion of the dealing and prior to leaving the BSC | ☐ |
| Procedures for the use of sharps, if applicable, including circumstances in which sharps to be used | ☐ |
| Procedures for movement of equipment out of facility, including its decontamination | ☐ |
| Procedures and validation for use of decontamination chamber if present | ☐ |
| Procedures for decontamination of GM organisms, including use of the autoclave | ☐ |
| Procedures and triggers for the gaseous decontamination | ☐ |
| Procedures for disposal of waste and effluent, including their transport | ☐ |
| Procedures for transporting viable GMOs inside the facility, including transport of GMOs for storage | ☐ |
| Procedures for the transport of viable material outside the facility (e.g. to another PC3 facility) | ☐ |
| Circumstances or events which must be reported to the Regulator | ☐ |
| Contingency plans for: |  |
| Spills of GMOs inside and outside the facility | ☐ |
| Accidental exposure/medical emergencies to the GMOs | ☐ |
| Escape of animals or invertebrates containing GMOs within or outside the facility, if applicable | ☐ |
| Alarms for fire or loss of pressure | ☐ |
| Release of GMOs from the facility (e.g., loss, theft or unintentional release) | ☐ |
| Failure of power or ventilation systems | ☐ |
| Failure of autoclaves and other decontamination systems | ☐ |
| Failure of BSCs and other aerosol containment equipment | ☐ |
| Natural disasters and fire | ☐ |
| Medical emergencies to persons | ☐ |
| Security threats | ☐ |

## Record keeping

Required to keep for	*at least 3 years of records		** at least 12 months of records

| Conditions | Item | Frequency | Completed | Comments |
| --- | --- | --- | --- | --- |
| C14, C15, C22 | Inspection of facility | Annually* | ☐ |  |
| C21 | Pest prevention strategy | Ongoing* | ☐ |  |
| C26 C28 C31 | BSC | Annually* | ☐ |  |
| C26 C28 C31 | IVCs & change stations | Annually* | ☐ |  |
| C26 C33 C34 | Autoclave (maintain) | Annually* | ☐ |  |
| C32 | Autoclave (validate) | Quarterly** | ☐ |  |
| C26 C33 C34 | LWTS | Annually* | ☐ |  |
| C36 | Backflow prevention device | Annually* | ☐ |  |
| C37 C38 | Ventilation | Annually* | ☐ |  |
| C68 | Gaseous decontamination | As required per C69 | ☐ |  |
| C72 | List of all GM organisms in the facility | As required per C72 | ☐ |  |
| C77-81 | Training/ Update training records | As required per C77-81* | ☐ |  |
| C83 | List of persons authorised to enter the facility | As required per C83 | ☐ |  |
| C84 | Facility manual review | Annually* | ☐ |  |

## Other

|  | Section | Requirement | Comments |
| --- | --- | --- | --- |
| C2 | Annual shutdown | Facility inspected after annual shutdown, if applicable ☐ |  |
| C11 | Work not including | Dealings above PC3 containment ☐ Housing of plants and aquatic animals beyond minimum time ☐ |  |
| C77 to C80 | Training | Authorised persons re-trained on last updated procedures? Training records updated ☐ Training of people who are authorised but not undertaking dealings ☐ e.g. tradie |  |
| C85 | Health monitoring | Procedures developed to: record accidents and potential exposure to GM human pathogens ☐ respond to exposure to GM human pathogens ☐ |  |
| C4 C5 | Authorised entry | Authority to exclude people from the facility ☐ Entry is restricted to authorised people ☐ |  |
| C21 | Pest Program | Program is proactive ☐ reactive ☐ Frequency |  |

# Work practices

|  | Section | Requirement | Comments |
| --- | --- | --- | --- |
| C47 C48 | PPE worn | lab coat ☐ closed footwear ☐ eye protection ☐ gloves ☐ skin abrasions covered ☐ Long hair tied/covered ☐ Optional Mask ☐ respirator ☐ face shield ☐ PAPR ☐ booties ☐ double gloves ☐ |  |
| R16 | PPE storage | Dedicated hanging/storage for PPE ☐ Reusable PPE are stored in a manner preventing cross contamination ☐ |  |
| C40 C41C42 | Personal effects | Food or drink not brought into/stored in the facility ☐ Personal items not taken into the facility ☐ Essential items decontaminated prior to exit ☐ |  |
| C74 | Transport of GMOs | Viable GMOs removed from the facility ☐ If yes, purpose of transport is: - Transport to other PC3 or PC4 facility ☐ - Export ☐ - Written permission from the Regulator ☐ Transport conducted in accordance with the Guidelines for the Transport, Storage and Disposal of GMOs ☐ |  |
| C52 C54 | Aerosol containment | Aerosol generating dealings with GMOs in BSC ☐ Centrifugation of GMOs in sealed containers ☐ |  |
| C55 | Sharps | Not used in direct connection with GMOs unless no alternatives are available ☐ |  |
| C23 C76 | Storage of GMOs | within work area only ☐ Triple contained ☐ | Exemptions? Alternate site inspected? |

# Inspection of the physical facility

## General facility

|  | Section | Requirement | Comments |
| --- | --- | --- | --- |
| R2 R3 | Facility Boundary | Facility sealable for gaseous decontamination ☐ Windows, if present, closed and sealed ☐ |  |
| C22 | Structural changes | Any since last inspection ☐ |  |
| C13 | Containment | Facility not used as a primary containment facility ☐ |  |
| R6 | Lifts | No lifts present ☐ |  |
| R8 | Pass through boxes /barrier autoclave | Type: Doors unable to be opened simultaneously ☐ Autoclave outer door opens only after successful decontamination ☐ |  |
| R10 R11 | Openings fitted with screens | Openings in walls, ceiling, or roof are screened ☐ Screens are: fixed and sealed against mounting ☐ made of appropriate material ☐ Small enough to prevent movement of invertebrates or animal ☐ |  |
| R5 | Penetrations | Sealed ☐ |  |
| R17 | Communication equipment | Primary equipment ☐ Alternative independent back-up system ☐ |  |
| R31 | Eyewash | Plumbed ☐ Single use ☐ |  |
| R13 R14 | Surfaces Fixtures/fittings | Surfaces are: smooth, impervious to water and easily cleanable ☐ resistant to cleaning agent and decontaminants ☐ Open spaces between and under fixtures, fittings and equipment accessible for decontamination ☐ |  |
| R9 | Decontamination Chamber | Present ☐ Doors unable to be opened simultaneously ☐ |  |

## Entry, airlock, anteroom

|  | Section | Requirement | Comments |
| --- | --- | --- | --- |
| R38 | Inward airflow | Air pressure at the entrance of the facility or each work area ☐ |  |
| R39 R40 |  | During the normal operation, work area at least 50 Pa below the pressure outside the facility ☐ and When either airlock door is opened, air pressure in all work area remains at least25 Pa below the pressure outside the facility ☐ |  |
| C16 | Signage on door | OGTR sign ☐ Biohazard sign ☐ Emergency contact numbers ☐ |  |
| C43 C44 | Facility access | Outer facility door lockable when unoccupied ☐ |  |
| C45 C46 |  | Entry and exit through airlock ☐ Emergency exits only used in an emergency if present ☐ |  |
| R7 | Airlock doors | Self-closing ☐ Sealed ☐ Outer door has mechanism to restrict entry ☐ Doors unable to be opened simultaneously ☐ |  |
| C18 | Airlock | No GMO dealings other than transport ☐ No used PPE ☐ No handwash basins ☐ No equipment ☐ |  |
| C19 | Anteroom | No GMO dealings other than transport (or autoclaving) ☐ No used PPE ☐ No handwash basins ☐ No equipment ☐ |  |

## Work area

|  | Section | Requirement | Comments |
| --- | --- | --- | --- |
| C17 | Emergency contacts | Visible in each work area ☐ |  |
| R30 | Hand decontamination | Present at or near the exit of work area ☐ Hands-free operation ☐ |  |
| R15 | GM dealing in multi room facility | Areas where GMO dealings are conducted are clearly identified ☐ |  |
| C73 | GMO labelling | Containers labelled to indicate they contains GMOs ☐ |  |

## Aerosol containment

|  | Section | Requirement | Comments |
| --- | --- | --- | --- |
| R20 | BSC | Number & type |  |
|  |  | Powered by uninterruptable power supply ☐ |  |
| R21 C28 C29 C30 C52 |  | Decontaminated before servicing ☐ Tested annually and after relocation ☐ HEPA filter integrity testing ☐ Result and date of the next test affixed to BSCs ☐ If defective, labelled appropriately ☐ |  |
|  |  | Used for all GMO dealings producing aerosols ☐ | Any exceptions? |
| C28b C29 C30 | IVCs and other aerosol containment devices approved in writing by the Regulator | Device type: Decontaminated before servicing ☐ Tested annually ☐ HEPA filter integrity testing ☐ Result and date of the next test affixed to equipment ☐ If defective, labelled appropriately ☐ |  |

## Decontamination

|  | Section | Requirement | Comments |
| --- | --- | --- | --- |
| R12 | Liquid effluent | Decontaminated prior to discharge from sink/floor drainage exit ☐ |  |
| C59 | GMOs/organisms infected with GMOs | Rendered non-viable before removal ☐ |  |
| C60 | Items removed from facility | Decontaminated prior to removal ☐ |  |
| C57 | Work surfaces | immediately after use ☐ at end of day ☐ |  |
| C49C61 | PPE | Removed when exiting work area ☐ Removed prior to entering the airlock ☐ Decontaminated before reuse or disposal ☐ |  |
| C25C58 | Equipment | Equipment decontaminated before maintenance ☐ |  |
| C63 | Other methods | UV / Irradiation ☐ Approved in writing by the Regulator ☐ |  |
| C70 | Spills | Spills kit present ☐ |  |

### Chemical decontamination

|  | Section | Requirement | Comments |
| --- | --- | --- | --- |
| C20 C67 | Chemical decontaminants | Available supply of disinfectants ☐ Concentration and exposure time effective against the GMOs being decontaminated ☐ Users aware of the minimum exposure time required for effective decontamination ☐ | Present in all work areas |
|  |  | Labelled (contents, concentration, expiry date) ☐ |  |

### Gaseous decontamination

|  | Section | Requirement | Comments |
| --- | --- | --- | --- |
| C68 | Gaseous decontamination | Chemical used for gaseous decontamination: |  |
|  |  | Has the chemical been validated as effective against the GMOs ☐ |  |
|  |  | Use of appropriate indicators for the validation of the decontamination ☐ |  |

### Heat decontamination

|  | Section | Requirement | Comments |
| --- | --- | --- | --- |
| R22 | Autoclave | present ☐ Not in airlock ☐ |  |
| C65 |  | Type: barrier ☐ free-standing ☐ Doors unable to be opened simultaneously ☐ |  |
| C32 |  | Physical parameters validated at least quarterly ☐ Via: Independent thermocouples ☐ or Chemical indicators which use a combination of moisture, heat, and time ☐ or Other ☐ |  |
| C66 |  | Combination of parameters (pressure, temperature, time) validated as effective against the organism being decontaminated or suitable surrogate ☐ |  |
| C33 |  | Annual testing and maintenance completed ☐, including: Calibration of temperature and pressure measuring probes ☐ Testing and confirmation that all parameters of the system (temperature, time, pressure and flowrate) are operating within the specified limits ☐ Testing of all safety and relief components, including the autoclave safety valves ☐ and Maintenance of the equipment to ensure effective operating conditions ☐ |  |
| C34 C35 |  | Result and date of the next test affixed to equipment ☐ If defective, labelled appropriately ☐ |  |
| C64 |  | Loads packed such that steam can be contact all material ☐ |  |
|  |  | Measures to differentiate loads pre and post decontamination | autoclave tape ☐ |
|  |  | Air, steam & liquid is filtered / decontaminated prior to discharge ☐ |  |
| R24 C33 C34 C60 C71 | LWTS | Present ☐ Fully enclosed system ☐ Pipes clearly identifiable ☐ Leak detection present ☐ or double skin ☐ Vents to pipes, tanks filtered ☐ Restricted access to LWTS room ☐ Validated to ensure that it can achieve the parameters required to inactivate the most resistant organism used in the facility ☐ |  |
|  |  | Annual testing and maintenance completed ☐, including: Calibration of temperature and pressure measuring probes ☐ Testing and confirmation that all parameters of the system (temperature, time, pressure and flowrate) are operating within the specified limits ☐ Testing of all safety and relief components, including the autoclave safety valves ☐ and Maintenance of the equipment to ensure effective operating conditions ☐ |  |

## Ventilation system

|  | Section | Requirement | Comments |
| --- | --- | --- | --- |
| R18 R19 | Room access | Access to ventilation equipment restricted ☐ Ducts, and access panels clearly identifiable ☐ |  |
| R32 R36 | Exhaust and supply system | The facility fitted with a ventilation system establishing negative air pressure in the facility and a directional airflow into the work area(s) ☐ Exhaust & supply system interlocked ☐ |  |
|  |  | Exhaust fitted with: At least one HEPA filter through which all exhaust air is filtered ☐ Pre filter upstream ☐ Exhaust fan downstream of the exhaust HEPA ☐ Gas-tight air ducts between the facility and the HEPA filter housing ☐ |  |
| R33 R34 R37 | Exhaust housing | HEPA mounted in gas-tight housing with sealed access doors ☐ |  |
|  |  | Design and location of the filter housing allow for: Easy access ☐ Secure and damage free handling of the HEPA ☐ In situ gaseous decontamination ☐ Maintenance, replacement and integrity testing of the HEPA filter ☐ If gaseous decon performed on HEPA filters, Gas tight isolating valve present on: Air inlet duct ☐ Upstream port ☐ and Downstream port If pressure across HEPA filter measured, 0.2µm hydrophobic membrane filter present on facility side of HEPA ☐ |  |
| C27 R41 R42 | Ventilation system malfunction and alarm | Facility equipped with an audible alarm ☐ Alarm activated when pressure in the work area(s) or airlock deviates from the set point by 15 Pa for more than 2 minutes ☐ In the event of failure or malfunction of the ventilation system: Use emergency stop button ☐ have procedures ☐ Notification to OGTR ☐ |  |
| R35 C25 C37 | Testing of the Ventilation system | Decontamination of the HEPA housing prior to testing or repair ☐ - Annual testing and maintenance of the facility ventilation system including: - Testing of the pressure differentials ☐ - Checking directional airflow ☐ - Testing of the alarm ☐ - Testing of components of the ventilation system in a system failure scenario ☐ - Calibration of transducers fitted to the air-handling system and validation of air-handling performance ☐ - Calibration of pressure gauges ☐ - The air handling control system ☐ - If applicable, the building management system ☐ and - Integrity testing of all HEPA ☐ |  |

## Backflow prevention devices

|  | Section | Requirement | Comments |
| --- | --- | --- | --- |
| R25 R26 R27 C36 | Backflow prevention | Water supply, connected directly or indirectly to any part of a water service, protected against backflow ☐ If no backflow prevention devices, documentation on risk assessment/risk management plan available ☐ If defective, labelled appropriately ☐ |  |
| R28 | Vacuum system | Present ☐ if so, fitted with: 0.2 µm hydrophobic membrane filters and liquid disinfectant traps ☐ If central vacuum system, filter and trap on the facility side of the vacuum line ☐ If portable pump, filter and trap located before the pump ☐ |  |
| R29 | Piped gas | Present in a BSC ☐ if so, Reverse flow prevention on outlets in BSC ☐ or fitted with ≤0.2µm hydrophobic membrane filter ☐ |  |

## Animal work area

|  | Section | Requirement | Comments |
| --- | --- | --- | --- |
| Ani-C1 Ani-C9 | Signs | Sign on outer door with special entry conditions/ animals present ☐ Documented system to account for all animals ☐ |  |
| Ani-C6 | Dealings generating aerosols | Conducted in BSC ☐ or in aerosol containment equipment approved in writing by the Regulator (e.g., Inhalation exposure system) ☐ |  |
| Ani-R1 Ani-C4 Ani-C5 Ani-C7 Ani-C8 | Cages/enclosure | Cages/enclosures devices appropriate for the animals being housed present (all life stages) ☐ Animal infected with aerosol transmissible GMOs housed in sealed cages/enclosures fitted with exhaust HEPA filters (IVCs or ventilated enclosure HEPA filtered) ☐ Cage is labelled to indicate it contains a GMO ☐ and the number of animals within the cage/enclosure ☐ |  |
| A-R2 | Containers used for Euthanising | Container used for euthanising animal containing GM micro-organisms requiring PC3 containment and capable of being shed are: sealed ☐ and fitted with a gas inlet filtered with a 0.2 µm hydrophobic membrane ☐ |  |
| Ani-C11 Ani-C12 | Decontamination | Animal carcasses decontaminated prior to removal from the facility ☐ Equipment, cages, bedding and animal waste decontaminated prior to removal from the facility ☐ |  |

## Invertebrate work area

|  | Section | Requirement | Comments |
| --- | --- | --- | --- |
| Inv-C1 Inv-C7 | Signs | Sign on outer door with special entry conditions/ invertebrates present ☐ Documented system to account for all invertebrates ☐ |  |
| Inv-R1 Inv-R2 Inv-R3 | Anteroom | Anteroom present for entry to the room(s) housing invertebrates ☐ Anteroom doors self-closing doors ☐ No dealings conducted within the anteroom ☐ Mechanism to prevent invertebrates from traversing anteroom/airlock boundary ☐ and capable to be maintained in darkened/ lit state ☐ Means present to inspect and remove invertebrates on persons exiting the anteroom ☐ |  |
| Inv-C5 Inv-C6 | Cages/enclosure | Cages/enclosures devices appropriate for the animals being housed present (all life stages) ☐ Cage is labelled to indicate it contains a GMO ☐ and the number of animals within the cage/enclosure ☐ |  |
| Inv-C10 | Aerosols | Dealings with aerosol transmissible GMOs dealings conducted in BSC ☐ |  |
| Inv-C11 Inv-C12 | Decontamination | Invertebrates euthanised prior to removal from the facility ☐ Cage/containers decontaminated prior to removal from the facility ☐ |  |
| Inv-C9 | Invertebrate trap | Room fitted with a suitable invertebrate trap ☐ |  |

33